{"text_format":"1","workbook":"xlsx","n_sheets":2,"sheets":["USFSeawallTestCO2comb.dat","Calibration"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"USFSeawallTestCO2comb.dat","index":0,"n_rows":25,"n_cols":16,"rows_kept":25,"cut":false,"hidden":false,"merged":[]}
[null,null,null,"ZERO_CNTS","ADC_CNTS","CO2","T_IRGA(AVG)","HUMIDITY(INT)","T_HUM","PRES(IRGA)","T_IRGA(DET)","T_IRGA(SRC)","DATETAG","TIMETAG2","Linearized correct zero reading","Recalcualted CO2"]
["April 9th, 2013",null,null,42640,40037,345.46300000000002,55.1,27.969000000000001,28.08,1084,54.8,55.4,"2013-099","12:30:01.570",44316,{"$f":"calcCO2(O2,E2,H2,I2,J2)","v":625.98729963313997}]
["Zero count was 44316 on February 5th and the data was a great match to the calibration gases, this was the last known good zero before deployment",null,null,42640,40037,345.44799999999998,55.1,27.969000000000001,28.08,1084,54.7,55.5,"2013-099","12:30:03.106",{"$f":"O2+0.00521897427","v":44316.005218974271},{"$f":"calcCO2(O3,E3,H3,I3,J3)","v":625.9882317039926}]
[null,null,null,42640,40037,345.43700000000001,55.1,27.969000000000001,28.08,1084,54.8,55.4,"2013-099","12:30:04.707",{"$f":"O3+0.00521897427","v":44316.010437948542},625.98916377523312]
[null,null,null,42640,40037,345.43,55.1,27.969000000000001,28.08,1084,54.7,55.5,"2013-099","12:30:06.308",44316.015656922813,625.9900958468603]
[null,null,null,42640,40037,345.42399999999998,55.1,27.969000000000001,28.08,1084,54.8,55.3,"2013-099","12:30:07.909",44316.020875897084,625.99102791887447]
[null,null,null,42640,40039,345.33499999999998,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:09.510",44316.026094871355,625.58553948734971]
[null,null,null,42640,40039,345.26900000000001,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:11.111",44316.031313845626,625.58647132386773]
[null,null,null,42640,40038,345.26,55.1,28.038,28.12,1084,54.8,55.4,"2013-099","12:30:12.711",44316.036532819897,625.78503754731412]
[null,null,null,42640,40037,345.29399999999998,55.1,28.038,28.12,1084,54.8,55.4,"2013-099","12:30:14.311",44316.041751794168,625.98363126796698]
[null,null,null,42640,40038,345.28,55.1,27.969000000000001,28.08,1084,54.7,55.5,"2013-099","12:30:15.912",44316.046970768439,625.79802099196525]
[null,null,null,42640,40038,345.26900000000001,55.1,27.969000000000001,28.08,1084,54.8,55.4,"2013-099","12:30:17.513",44316.05218974271,625.79895296091627]
[null,null,null,42640,40037,345.298,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:19.114",44316.057408716981,625.98642741445678]
[null,null,null,42640,40039,345.238,55.1,28.038,28.12,1084,54.8,55.4,"2013-099","12:30:20.714",44316.062627691252,625.59206235110548]
[null,null,null,42640,40039,345.19299999999998,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:22.315",44316.067846665523,625.59299419033323]
[null,null,null,42640,40039,345.16,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:23.916",44316.073065639794,625.59392602994831]
[null,null,null,42640,40039,345.13400000000001,55.1,28.038,28.12,1084,54.8,55.4,"2013-099","12:30:25.517",44316.078284614065,625.59485786994981]
[null,null,null,42640,40040,345.07299999999998,55.1,28.038,28.12,1084,54.8,55.5,"2013-099","12:30:27.116",44316.083503588336,625.39818165938573]
[null,null,null,42640,40040,345.02499999999998,55.1,28.106000000000002,28.16,1084,54.8,55.4,"2013-099","12:30:28.718",44316.088722562607,625.3882635324394]
[null,null,null,42640,40039,345.03300000000002,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:30.317",44316.093941536878,625.59765339227965]
[null,null,null,42640,40039,345.03800000000001,55.1,28.038,28.12,1084,54.7,55.4,"2013-099","12:30:31.917",44316.099160511149,625.59858523383002]
[null,null,null,42640,40040,344.99900000000002,55.1,28.106000000000002,28.16,1084,54.8,55.4,"2013-099","12:30:33.520",44316.104379485419,625.39105866557361]
[null,null,null,42640,40162,329.75099999999998,55.1,28.411999999999999,28.96,1072,54.8,55.4,"2013-099","13:30:01.572",44316.10959845969,610.77040161860066]
[null,null,null,42640,40161,329.88499999999999,55.1,28.344000000000001,28.92,1072,54.7,55.4,"2013-099","13:30:03.153",44316.114817433961,610.97883925894382]
[null,null,null,42640,40161,329.99900000000002,55.1,28.161000000000001,28.92,1072,54.8,55.5,"2013-099","13:30:04.754",44316.120036408232,611.02393003230327]

{"sheet":"Calibration","index":1,"n_rows":21,"n_cols":6,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
[8.0126290000000004,"Cscale",null,"s/n 4081 cal from 8/30/2012"]
[0.1,"C1a"]
[624.03689999999995,"C1b"]
[316.04070000000002,"C1c"]
[67.949259999999995,"C1d"]
[1.1100000000000001,"Cbreak1",null,null,"Notes:"]
[-531.17550000000006,"C2a",null,null,"User Defined Function calcCO2(Zero, A/D, H2O, H2OTemp, Press) performs all the calculations that the SBA-5 firmware performs"]
[1717.078,"C2b",null,null,"Uses calibration data on this sheet - do not rename this sheet, or move the locations of any data, or add a row."]
[-412.36579999999998,"C2c",null,null,"Pressure compensation factors in blue do not change from one SBA-5 to another, but do change for a high range SBA-5 (with a short IRGA)."]
[225.49760000000001,"C2d",null,null,"If no humidity sensor is present, pass 0 for H2O and 0 for H2OTemp in function call."]
[3.23,"Cbreak2"]
[-176302.7,"C3a"]
[151143.5,"C3b"]
[-42654.99,"C3c"]
[4197.1689999999999,"C3d"]
[6.7039589999999996E-2,"cha",null,null,"Pressure Compensation"]
[3.9226190000000001E-4,"chb",null,null,"Standard","High Range"]
[-0.40012560000000003,"Pa",null,null,-0.40012560000000003,-0.55785079999999998]
[1.3976539999999999E-3,"Pb",null,null,1.4E-3,1.5653080000000001E-3]
[-2.1598729999999999E-5,"Pc",null,null,-2.16E-5,-2.8200000000000001E-6]
[2.145194E-8,"Pd",null,null,2.1500000000000001E-8,2.7000000000000002E-9]